## Sheet1
Sample | [9 empty] | 23 Na [ He Gas ] |  |  | 24 Mg [ He Gas ] |  |  | 27 Al [ He Gas ] |  |  | 39 K [ He Gas ] |  |  | 44 Ca [ He Gas ] |  |  | 52 Cr [ He Gas ] |  |  | 55 Mn [ He Gas ] |  |  | 56 Fe [ He Gas ] |  |  | 60 Ni [ He Gas ] |  |  | 63 Cu [ He Gas ] |  |  | 66 Zn [ He Gas ] |  |  | 75 As [ He Gas ] |  |  | 78 Se [ He Gas ] |  |  | 82 Se [ He Gas ] |  |  | 95 Mo [ He Gas ] |  |  | 137 Ba [ He Gas ] |  |  | 208 Pb [ He Gas ] |  |  | 7 Li ( ISTD ) [ He Gas ] |  | 45 Sc ( ISTD ) [ He Gas ] |  | 72 Ge ( ISTD ) [ He Gas ] |  | 103 Rh ( ISTD ) [ He Gas ] |  | 115 In ( ISTD ) [ He Gas ] |  | 159 Tb ( ISTD ) [ He Gas ] |  | 175 Lu ( ISTD ) [ He Gas ]
| Rjct | Data File | Acq. Date-Time | Type | Level | Sample Name | Total Dil. | Comment | Vial Number | Conc. [ ppb ] | Conc. RSD | Meas. Conc. [ ppb ] | Conc. [ ppb ] | Conc. RSD | Meas. Conc. [ ppb ] | Conc. [ ppb ] | Conc. RSD | Meas. Conc. [ ppb ] | Conc. [ ppb ] | Conc. RSD | Meas. Conc. [ ppb ] | Conc. [ ppb ] | Conc. RSD | Meas. Conc. [ ppb ] | Conc. [ ppb ] | Conc. RSD | Meas. Conc. [ ppb ] | Conc. [ ppb ] | Conc. RSD | Meas. Conc. [ ppb ] | Conc. [ ppb ] | Conc. RSD | Meas. Conc. [ ppb ] | Conc. [ ppb ] | Conc. RSD | Meas. Conc. [ ppb ] | Conc. [ ppb ] | Conc. RSD | Meas. Conc. [ ppb ] | Conc. [ ppb ] | Conc. RSD | Meas. Conc. [ ppb ] | Conc. [ ppb ] | Conc. RSD | Meas. Conc. [ ppb ] | Conc. [ ppb ] | Conc. RSD | Meas. Conc. [ ppb ] | Conc. [ ppb ] | Conc. RSD | Meas. Conc. [ ppb ] | Conc. [ ppb ] | Conc. RSD | Meas. Conc. [ ppb ] | Conc. [ ppb ] | Conc. RSD | Meas. Conc. [ ppb ] | Conc. [ ppb ] | Conc. RSD | Meas. Conc. [ ppb ] | CPS | CPS RSD | CPS | CPS RSD | CPS | CPS RSD | CPS | CPS RSD | CPS | CPS RSD | CPS | CPS RSD | CPS | CPS RSD
| FALSE | 001CALB.d | 2020-07-16 09:45 | CalBlk | 1 | Calibration Blank | 1 |  | 1101 | 0 | N/A | 0 |  |  |  | 0 | N/A | 0 | 0 | N/A | 0 | 0 | N/A | 0 | 0 | N/A | 0 | 0 | N/A | 0 | 0 | N/A | 0 | 0 | N/A | 0 | 0 | N/A | 0 | 0 | N/A | 0 | 0 | N/A | 0 | 0 | N/A | 0 | 0 | N/A | 0 | 0 | N/A | 0 | 0 | N/A | 0 | 0 | N/A | 0 | 185.56 | 2.0722 | 246915.7767 | 0.34558 | 137013.3533 | 0.69574 | 1949622.6333 | 0.38694 | 1371155.3233 | 0.23592 | 2707504.0567 | 0.78869 | 2206373.5267 | 1.1531
| FALSE | 002CALS.d | 2020-07-16 09:49 | CalStd | 2 | 0.5 ppb / 25 ppb | 1 |  | 1102 | 23.0575 | 2.1151 | 23.0575 |  |  |  | <0.000 | N/A | <0.000 | 24.197 | 4.8807 | 24.197 | 20.5295 | 3.6042 | 20.5295 | 0.50384 | 0.97091 | 0.50384 | 0.49959 | 0.93598 | 0.49959 | 25.67 | 0.77988 | 25.67 | 0.51809 | 2.0631 | 0.51809 | 0.53943 | 2.341 | 0.53943 | 25.7324 | 0.50148 | 25.7324 | 0.50965 | 1.1769 | 0.50965 | 0.54714 | 3.2346 | 0.54714 | 0.42083 | 11.9573 | 0.42083 | 0.48196 | 3.3827 | 0.48196 | 0.4965 | 2.9096 | 0.4965 | 0.50034 | 0.69904 | 0.50034 | 161.1133 | 20.8278 | 244825.6333 | 0.99161 | 136474.1567 | 0.97974 | 1930093.1167 | 1.6681 | 1367399.0833 | 0.88573 | 2714688.3633 | 0.79699 | 2191215.2633 | 0.76244
| FALSE | 003CALS.d | 2020-07-16 09:53 | CalStd | 3 | 1.0 ppb / 100 ppb | 1 |  | 1103 | 99.1763 | 1.0691 | 99.1763 |  |  |  | <0.000 | N/A | <0.000 | 100.1186 | 0.80571 | 100.1186 | 93.9549 | 1.2794 | 93.9549 | 1.0265 | 1.0855 | 1.0265 | 1.0006 | 1.4621 | 1.0006 | 103.3664 | 1.1097 | 103.3664 | 1.0465 | 4.114 | 1.0465 | 1.0707 | 0.74862 | 1.0707 | 104.8109 | 0.95236 | 104.8109 | 1.0284 | 1.2542 | 1.0284 | 1.0429 | 5.6942 | 1.0429 | 1.0724 | 13.2756 | 1.0724 | 0.97393 | 1.6636 | 0.97393 | 1.0184 | 1.2183 | 1.0184 | 1.0288 | 0.31747 | 1.0288 | 191.1133 | 13.5472 | 246131.0133 | 0.86578 | 135254.1533 | 0.79026 | 1954115.34 | 0.83374 | 1369537.4033 | 0.5906 | 2717905.6533 | 0.49814 | 2187823.59 | 0.58128
| FALSE | 004CALS.d | 2020-07-16 09:56 | CalStd | 4 | 5.0 ppb / 250 ppb | 1 |  | 1104 | 259.6947 | 0.81581 | 259.6947 |  |  |  | 3.7233 | 2.2906 | 3.7233 | 261.1665 | 0.78135 | 261.1665 | 243.747 | 0.21196 | 243.747 | 5.1279 | 0.26326 | 5.1279 | 5.1165 | 0.44348 | 5.1165 | 257.5688 | 0.97745 | 257.5688 | 5.3111 | 0.78344 | 5.3111 | 5.4213 | 0.64993 | 5.4213 | 265.9232 | 0.078932 | 265.9232 | 5.1704 | 0.27661 | 5.1704 | 5.1263 | 2.1844 | 5.1263 | 5.1429 | 5.1193 | 5.1429 | 4.9905 | 0.36628 | 4.9905 | 5.1109 | 1.4759 | 5.1109 | 5.2408 | 1.686 | 5.2408 | 151.1133 | 6.7395 | 246369.72 | 0.39139 | 134891.8767 | 0.62085 | 1911639.3 | 0.96153 | 1362701.4333 | 1.0403 | 2725412.74 | 0.74165 | 2188388.5933 | 1.5775
| TRUE | 005CALS.d | 2020-07-16 10:00 | CalStd | 5 | 10 ppb / 500 ppb | 1 |  | 1105
| FALSE | 006CALS.d | 2020-07-16 10:04 | CalStd | 6 | 20 ppb / 1000 ppb | 1 |  | 1106 | 979.1936 | 0.74201 | 979.1936 |  |  |  | 18.9876 | 0.75701 | 18.9876 | 984.3899 | 0.49896 | 984.3899 | 961.5441 | 0.53122 | 961.5441 | 20.2476 | 0.9454 | 20.2476 | 20.2814 | 0.54789 | 20.2814 | 981.8686 | 0.7341 | 981.8686 | 20.6028 | 0.69639 | 20.6028 | 21.101 | 0.44857 | 21.101 | 1004.9569 | 0.35265 | 1004.9569 | 20.2875 | 0.044108 | 20.2875 | 20.3588 | 1.1352 | 20.3588 | 20.2481 | 0.3307 | 20.2481 | 19.3865 | 1.1455 | 19.3865 | 20.1331 | 1.0397 | 20.1331 | 20.8033 | 0.45449 | 20.8033 | 164.45 | 22.2347 | 243250.3367 | 0.45065 | 134556.5133 | 0.51655 | 1931893.6733 | 1.3646 | 1347565.6733 | 0.66092 | 2713274.2667 | 0.44205 | 2164653.04 | 1.0371
| FALSE | 007CALS.d | 2020-07-16 10:08 | CalStd | 5 | 10 ppb / 500 ppb | 1 |  | 1105 | 520.9787 | 1.0454 | 520.9787 |  |  |  | 9.182 | 1.2276 | 9.182 | 525.8539 | 2.2076 | 525.8539 | 500.5654 | 0.0079361 | 500.5654 | 10.4235 | 0.73298 | 10.4235 | 10.4029 | 0.59893 | 10.4029 | 522.5316 | 0.92309 | 522.5316 | 10.7831 | 0.58408 | 10.7831 | 11.0591 | 0.64525 | 11.0591 | 542.358 | 0.83953 | 542.358 | 10.5468 | 0.39374 | 10.5468 | 10.5773 | 1.8214 | 10.5773 | 10.6357 | 2.2831 | 10.6357 | 10.1753 | 0.54459 | 10.1753 | 10.3832 | 0.24327 | 10.3832 | 10.6693 | 1.3612 | 10.6693 | 156.67 | 13.9504 | 243887.0133 | 0.10987 | 132791.43 | 0.27849 | 1914605.62 | 0.74073 | 1350063.1233 | 0.48631 | 2715818.71 | 0.31006 | 2171273.1067 | 1.2972
| FALSE | 008CALS.d | 2020-07-16 10:12 | CalStd | 7 | 50 ppb / 2500 ppb | 1 |  | 1107 | 2564.4241 | 0.90055 | 2564.4241 |  |  |  | 49.881 | 1.1377 | 49.881 | 2574.0676 | 0.46855 | 2574.0676 | 2511.6901 | 0.34216 | 2511.6901 | 51.9747 | 0.74195 | 51.9747 | 52.683 | 1.7953 | 52.683 | 2572.6865 | 1.3369 | 2572.6865 | 52.4524 | 0.77259 | 52.4524 | 54.4659 | 1.1363 | 54.4659 | 2601.8643 | 0.10508 | 2601.8643 | 51.594 | 1.2012 | 51.594 | 51.7017 | 1.4348 | 51.7017 | 51.7223 | 1.3549 | 51.7223 | 49.5909 | 0.49774 | 49.5909 | 51.5004 | 0.50406 | 51.5004 | 51.9072 | 0.71819 | 51.9072 | 135.56 | 7.5119 | 241665.8233 | 0.64971 | 132525.1833 | 0.83494 | 1916571.1733 | 0.78201 | 1328427.2633 | 0.32383 | 2650814.9633 | 0.31179 | 2178476.0933 | 0.62426
| FALSE | 009CALS.d | 2020-07-16 10:15 | CalStd | 8 | 100 ppb / 5000 ppb | 1 |  | 1201 | 5154.2783 | 1.0279 | 5154.2783 |  |  |  | 100.8133 | 0.58259 | 100.8133 | 5129.762 | 0.40679 | 5129.762 | 5115.0618 | 0.84405 | 5115.0618 | 101.3595 | 0.60466 | 101.3595 | 102.0052 | 0.70736 | 102.0052 | 5164.1342 | 0.8957 | 5164.1342 | 103.5017 | 0.58555 | 103.5017 | 104.0858 | 0.30356 | 104.0858 | 5206.9883 | 0.025786 | 5206.9883 | 101.431 | 0.42158 | 101.431 | 101.9532 | 0.56994 | 101.9532 | 101.1508 | 0.55908 | 101.1508 | 102.3303 | 0.21343 | 102.3303 | 101.2746 | 0.48302 | 101.2746 | 101.9782 | 0.44741 | 101.9782 | 184.4467 | 7.3027 | 237577.71 | 0.24058 | 131861.97 | 0.53179 | 1861093.6767 | 0.47061 | 1310498.05 | 0.61219 | 2648877.3967 | 0.36534 | 2160949.0133 | 0.51596
| FALSE | 010CALS.d | 2020-07-16 10:19 | CalStd | 9 | 200 ppb / 10000 ppb | 1 |  | 1202 | 9907.5524 | 1.4786 | 9907.5524 |  |  |  | 199.7972 | 0.71015 | 199.7972 | 9916.5901 | 0.47828 | 9916.5901 | 9943.5806 | 0.3325 | 9943.5806 | 198.7773 | 0.8071 | 198.7773 | 198.2755 | 0.95873 | 198.2755 | 9900.2233 | 0.36777 | 9900.2233 | 197.5286 | 1.936 | 197.5286 | 196.6671 | 1.3019 | 196.6671 | 9867.9782 | 0.35901 | 9867.9782 | 198.8255 | 0.5411 | 198.8255 | 198.5299 | 0.76875 | 198.5299 | 198.9335 | 1.4014 | 198.9335 | 198.9901 | 0.4994 | 198.9901 | 198.9523 | 1.2681 | 198.9523 | 198.4141 | 0.51608 | 198.4141 | 163.3367 | 15.9407 | 236364.7567 | 1.2114 | 132230.3567 | 0.50711 | 1858923.12 | 0.52855 | 1296207.39 | 0.36326 | 2672826.63 | 1.4493 | 2144433.25 | 0.7597
| TRUE | 011CALS.d | 2020-07-16 10:23 | CalStd | 3 | 100 ppb (Mg) | 1 |  | 2501
| TRUE | 012CALS.d | 2020-07-16 10:26 | CalStd | 5 | 500 ppb (Mg) | 1 |  | 2502
| TRUE | 013CALB.d | 2020-07-16 10:30 | CalBlk | 1 | Calibration Blank | 1 |  | 1101
| TRUE | 014CALS.d | 2020-07-16 10:33 | CalStd | 6 | 1000 ppb (Mg) | 1 |  | 2503
| TRUE | 015CALS.d | 2020-07-16 10:36 | CalStd | 8 | 5000 ppb (Mg) | 1 |  | 2504
| TRUE | 016CALS.d | 2020-07-16 10:40 | CalStd | 9 | 10000 ppb (Mg) | 1 |  | 2505
| FALSE | 017_ICV.d | 2020-07-16 10:43 | ICV |  | ICV 20 ppb / 1000 ppb | 1 |  | 1106 | 1003.709 | 1.079 | 1003.709 |  |  |  | 19.0618 | 1.4674 | 19.0618 | 989.7439 | 1.8319 | 989.7439 | 954.126 | 0.88719 | 954.126 | 20.1012 | 0.23916 | 20.1012 | 20.1334 | 0.15443 | 20.1334 | 1017.2882 | 0.46443 | 1017.2882 | 20.8906 | 0.4205 | 20.8906 | 21.3422 | 0.98091 | 21.3422 | 1033.1432 | 0.82083 | 1033.1432 | 20.4586 | 0.33087 | 20.4586 | 20.4373 | 2.3241 | 20.4373 | 20.4502 | 1.2029 | 20.4502 | 18.973 | 0.38962 | 18.973 | 19.6487 | 0.56928 | 19.6487 | 21.067 | 0.35182 | 21.067 | 170.0033 | 7.0715 | 239419.72 | 0.70003 | 130125.5467 | 0.1835 | 1926714.6433 | 0.92932 | 1383587.6067 | 0.81017 | 2693728.5733 | 0.71515 | 2186342.1367 | 0.47416
| FALSE | 018New2.d | 2020-07-16 10:47 | BLK |  | Cal. Blank QC | 1 |  | 1101 | <0.000 | N/A | <0.000 |  |  |  | <0.000 | N/A | <0.000 | <0.000 | N/A | <0.000 | <0.000 | N/A | <0.000 | 0.0024134 | 125.9235 | 0.0024134 | <0.000 | N/A | <0.000 | <0.000 | N/A | <0.000 | <0.000 | N/A | <0.000 | 0.0022602 | 16.8457 | 0.0022602 | 0.058258 | 4.7931 | 0.058258 | 0.0029543 | 34.5285 | 0.0029543 | <0.000 | N/A | <0.000 | <0.000 | N/A | <0.000 | 0.037562 | 3.0331 | 0.037562 | <0.000 | N/A | <0.000 | 0.21215 | 5.453 | 0.21215 | 160.0033 | 19.8734 | 236338.0967 | 0.15081 | 128860.6 | 0.18464 | 1919939.23 | 0.80325 | 1340771.5233 | 0.16773 | 2656503.6433 | 1.3758 | 2168311.3767 | 0.17539
| FALSE | 019SMPL.d | 2020-07-16 10:51 | QCS |  | QCS | 1 |  | 2101 | 2368.1745 | 1.2311 | 2368.1745 |  |  |  | 48.6125 | 1.1966 | 48.6125 | 2356.7804 | 0.54867 | 2356.7804 | 2312.6319 | 0.13789 | 2312.6319 | 49.7399 | 1.2608 | 49.7399 | 49.8439 | 2.216 | 49.8439 | 2366.2498 | 0.43454 | 2366.2498 | 48.9228 | 0.58374 | 48.9228 | 52.3966 | 1.2322 | 52.3966 | 2424.3134 | 0.7796 | 2424.3134 | 49.7842 | 0.49147 | 49.7842 | 48.271 | 0.59549 | 48.271 | 48.8276 | 0.38557 | 48.8276 | 45.2395 | 0.40188 | 45.2395 | 46.0556 | 0.89518 | 46.0556 | 48.6442 | 0.14102 | 48.6442 | 180.0067 | 22.4517 | 235416.4833 | 0.4178 | 129420.4633 | 0.06862 | 1892237.84 | 0.48767 | 1322157.4467 | 0.55996 | 2697182.53 | 1.0638 | 2176090.54 | 0.4359
| FALSE | 020SMPL.d | 2020-07-16 10:55 | LFB |  | LFB | 1 |  | 2102 | 2405.9479 | 0.11188 | 2405.9479 |  |  |  | 45.8582 | 0.32954 | 45.8582 | 2369.1197 | 0.44337 | 2369.1197 | 2294.5316 | 0.37423 | 2294.5316 | 48.1645 | 1.475 | 48.1645 | 47.836 | 0.39273 | 47.836 | 2395.2542 | 0.28555 | 2395.2542 | 48.1513 | 0.84088 | 48.1513 | 50.3266 | 0.79189 | 50.3266 | 2437.5098 | 0.90829 | 2437.5098 | 47.134 | 0.74296 | 47.134 | 47.1085 | 1.011 | 47.1085 | 47.3461 | 1.3586 | 47.3461 | 44.4393 | 0.42082 | 44.4393 | 45.6307 | 0.55995 | 45.6307 | 47.6402 | 0.14719 | 47.6402 | 182.2267 | 17.2895 | 234902.3867 | 0.31155 | 128996.8267 | 1.0728 | 1912704.9267 | 0.26157 | 1312467.5767 | 0.44141 | 2684944.13 | 0.38644 | 2168897.69 | 0.25481
| FALSE | 021SMPL.d | 2020-07-16 10:58 | LRB |  | LRB | 1 |  | 2103 | <0.000 | N/A | <0.000 |  |  |  | 0.093571 | 205.32 | 0.093571 | <0.000 | N/A | <0.000 | <0.000 | N/A | <0.000 | 0.0083875 | 65.1519 | 0.0083875 | <0.000 | N/A | <0.000 | <0.000 | N/A | <0.000 | 0.20909 | 15.7658 | 0.20909 | 0.032701 | 15.721 | 0.032701 | 0.30195 | 37.6711 | 0.30195 | 0.0062028 | 23.2033 | 0.0062028 | <0.000 | N/A | <0.000 | <0.000 | N/A | <0.000 | 0.092761 | 12.8271 | 0.092761 | <0.000 | N/A | <0.000 | 0.045268 | 11.7786 | 0.045268 | 173.34 | 5.769 | 237436.7633 | 9.1493 | 130461.4467 | 4.3248 | 1933331.1733 | 10.4138 | 1359894.4267 | 11.4694 | 2695028.5067 | 9.4288 | 2179339.9133 | 8.3193
| FALSE | 022SMPL.d | 2020-07-16 11:02 | Sample |  | 1 | 100 |  | 2104 | 71717.2714 | 1.0999 | 717.1727 |  |  |  | 196.3129 | 5.6682 | 1.9631 | 1656848.0843 | 0.33796 | 16568.4808 | 399066.6142 | 0.17536 | 3990.6661 | 23.8865 | 2.9687 | 0.23886 | 1135.1235 | 0.8018 | 11.3512 | 5408.9909 | 0.5066 | 54.0899 | 166.0177 | 0.72924 | 1.6602 | 145.0708 | 1.4536 | 1.4507 | 311.2252 | 0.74357 | 3.1123 | 5.983 | 5.3936 | 0.05983 | 46.0779 | 31.5396 | 0.46078 | 39.4275 | 6.685 | 0.39427 | 9.7894 | 6.1529 | 0.097894 | 706.6192 | 0.87903 | 7.0662 | 7.4454 | 1.7163 | 0.074454 | 203.3367 | 24.48 | 235307.6067 | 0.37924 | 129869.6333 | 0.31212 | 1889428.4667 | 1.4085 | 1303673.32 | 0.43342 | 2693555.1 | 0.89998 | 2188248.4533 | 1.2766
| FALSE | 023SMPL.d | 2020-07-16 11:06 | Sample |  | 2 | 100 |  | 2105 | 182279.2041 | 1.1204 | 1822.792 |  |  |  | 16431.5989 | 0.44928 | 164.316 | 1188482.5571 | 0.2208 | 11884.8256 | 1984536.6084 | 0.97618 | 19845.3661 | 47.7151 | 1.7312 | 0.47715 | 16128.206 | 0.87923 | 161.2821 | 11933.9648 | 1.205 | 119.3396 | 142.3822 | 0.29945 | 1.4238 | 53.7 | 0.9233 | 0.537 | 336.1694 | 1.5313 | 3.3617 | 27.8431 | 2.2985 | 0.27843 | 41.6097 | 18.4498 | 0.4161 | 40.027 | 38.3198 | 0.40027 | 18.4498 | 2.3676 | 0.1845 | 2211.6181 | 0.64948 | 22.1162 | 186.3363 | 0.48612 | 1.8634 | 288.8933 | 9.2541 | 239133.6533 | 0.40534 | 130809.22 | 0.61982 | 1903986.59 | 0.62796 | 1324616.5967 | 0.16304 | 2730596.21 | 0.13789 | 2196294.8467 | 1.0249
| FALSE | 024SMPL.d | 2020-07-16 11:09 | Sample |  | 3 | 100 |  | 2106 | 183565.3976 | 1.3393 | 1835.654 |  |  |  | 72684.9152 | 0.55457 | 726.8492 | 1146825.4065 | 0.33474 | 11468.2541 | 1920096.0068 | 1.5169 | 19200.9601 | 83.9199 | 1.0722 | 0.8392 | 18632.3193 | 1.3343 | 186.3232 | 60865.0127 | 0.24562 | 608.6501 | 145.0463 | 0.89648 | 1.4505 | 78.7173 | 0.56367 | 0.78717 | 407.3153 | 1.1816 | 4.0732 | 30.9253 | 3.5021 | 0.30925 | 43.5038 | 4.3157 | 0.43504 | 55.8896 | 25.3911 | 0.5589 | 20.6703 | 2.0706 | 0.2067 | 3008.1641 | 0.92303 | 30.0816 | 228.7027 | 0.88749 | 2.287 | 478.9033 | 7.7613 | 238416.1267 | 0.68339 | 132449.17 | 0.44079 | 1893950.9667 | 1.16 | 1325785.1733 | 0.22727 | 2753112.0467 | 1.0027 | 2235319.7033 | 0.51058
| FALSE | 025SMPL.d | 2020-07-16 11:13 | Sample |  | 4 | 100 |  | 2107 | 176205.8546 | 1.153 | 1762.0585 |  |  |  | <0.000 | N/A | <0.000 | 134622.8212 | 1.3444 | 1346.2282 | 27681.4037 | 1.1233 | 276.814 | 3.4528 | 3.9753 | 0.034528 | 135.5754 | 1.6042 | 1.3558 | 315.5176 | 2.1792 | 3.1552 | 12.9534 | 31.6118 | 0.12953 | 29.6248 | 2.3484 | 0.29625 | 165.0742 | 0.72314 | 1.6507 | 8.5997 | 2.3793 | 0.085997 | <0.000 | N/A | <0.000 | <0.000 | N/A | <0.000 | 6.1796 | 6.7957 | 0.061796 | 25.761 | 1.9076 | 0.25761 | 2.1305 | 8.3376 | 0.021305 | 156.67 | 11.8468 | 236384.33 | 0.54445 | 132952.41 | 0.061882 | 1924785.1333 | 0.37394 | 1338687.1033 | 0.41657 | 2717700.6567 | 1.0269 | 2184444.1467 | 1.5877
| FALSE | 026SMPL.d | 2020-07-16 11:17 | Sample |  | 5 | 100 |  | 2108 | 117956.6802 | 1.3914 | 1179.5668 |  |  |  | 10748.7569 | 0.36049 | 107.4876 | 402690.9948 | 0.43252 | 4026.9099 | 244332.5616 | 0.14637 | 2443.3256 | 15.69 | 2.0043 | 0.1569 | 4054.5136 | 0.397 | 40.5451 | 16056.0455 | 0.90095 | 160.5605 | 49.5887 | 1.5523 | 0.49589 | 27.7511 | 1.8573 | 0.27751 | <0.000 | N/A | <0.000 | 12.3428 | 2.4796 | 0.12343 | 6.584 | 58.3465 | 0.06584 | <0.000 | N/A | <0.000 | 6.0433 | 7.2466 | 0.060433 | 473.2542 | 1.3978 | 4.7325 | 9.772 | 1.5614 | 0.09772 | 197.7833 | 19.9655 | 239253.14 | 0.11201 | 131965.9433 | 0.15992 | 1924557.2167 | 0.51138 | 1338788.4467 | 0.3494 | 2742618.99 | 1.0628 | 2219665.12 | 0.81075
| FALSE | 027SMPL.d | 2020-07-16 11:20 | Sample |  | 6 | 100 |  | 2109 | 43967.9564 | 1.324 | 439.6796 |  |  |  | 128.9659 | 5.0061 | 1.2897 | 218183.0556 | 0.68822 | 2181.8306 | 125605.2211 | 0.94895 | 1256.0522 | 3.4773 | 3.8201 | 0.034773 | 619.7168 | 0.66883 | 6.1972 | 1851.7067 | 1.1624 | 18.5171 | 21.6481 | 8.1589 | 0.21648 | 36.9772 | 1.2143 | 0.36977 | <0.000 | N/A | <0.000 | 6.2514 | 4.6983 | 0.062514 | <0.000 | N/A | <0.000 | <0.000 | N/A | <0.000 | 6.3477 | 6.5467 | 0.063477 | 115.7615 | 4.9354 | 1.1576 | 3.0674 | 24.7877 | 0.030674 | 168.8933 | 7.4734 | 238897.24 | 1.0128 | 132006.41 | 0.84747 | 1930069.4367 | 0.86343 | 1331539.3767 | 0.51607 | 2721049.4767 | 0.49512 | 2209871.9267 | 0.97895
| FALSE | 028SMPL.d | 2020-07-16 11:24 | Sample |  | 1 | 1000 |  | 2110 | 64119.2438 | 1.0775 | 64.1192 |  |  |  | <0.000 | N/A | <0.000 | 1562738.0556 | 0.57376 | 1562.7381 | 366891.5086 | 0.76385 | 366.8915 | 29.7597 | 5.9313 | 0.02976 | 1085.902 | 1.467 | 1.0859 | 4695.0367 | 0.82002 | 4.695 | 146.0092 | 2.1912 | 0.14601 | 135.1938 | 1.9384 | 0.13519 | <0.000 | N/A | <0.000 | 6.6587 | 13.8236 | 0.0066587 | 19.8678 | 185.9565 | 0.019868 | 62.2738 | 263.7891 | 0.062274 | 9.2732 | 20.953 | 0.0092732 | 664.715 | 1.3105 | 0.66471 | 8.9301 | 21.249 | 0.0089301 | 163.3367 | 14.7168 | 236569.92 | 1.0953 | 132276.0833 | 0.7459 | 1941540.6167 | 1.2764 | 1330831.7667 | 0.18095 | 2723509.6833 | 0.6259 | 2214438.73 | 0.91366
| FALSE | 029SMPL.d | 2020-07-16 11:28 | Sample |  | 2 | 1000 |  | 2111 | 155012.9337 | 0.6381 | 155.0129 |  |  |  | 12541.6037 | 0.48258 | 12.5416 | 995889.2223 | 0.90492 | 995.8892 | 1605534.4119 | 0.35787 | 1605.5344 | 42.9758 | 1.6515 | 0.042976 | 13261.048 | 0.43845 | 13.261 | 9498.4132 | 1.4219 | 9.4984 | 98.5429 | 5.4693 | 0.098543 | 49.2337 | 4.9865 | 0.049234 | <0.000 | N/A | <0.000 | 20.696 | 3.5608 | 0.020696 | <0.000 | N/A | <0.000 | <0.000 | N/A | <0.000 | 17.1452 | 12.1892 | 0.017145 | 1858.1896 | 1.3563 | 1.8582 | 153.7912 | 1.7493 | 0.15379 | 166.67 | 7.2096 | 236469.3233 | 0.35515 | 130612.9767 | 0.75819 | 1923161.5167 | 1.1073 | 1334805.5633 | 0.0039672 | 2724498.0167 | 0.60739 | 2193596.93 | 0.76941
| FALSE | 030SMPL.d | 2020-07-16 11:31 | Sample |  | 3 | 1000 |  | 2112 | 166977.8749 | 0.38406 | 166.9779 |  |  |  | 58962.631 | 0.37186 | 58.9626 | 1011598.4539 | 0.81298 | 1011.5985 | 1652964.1193 | 0.45562 | 1652.9641 | 80.7937 | 6.6031 | 0.080794 | 16497.735 | 0.89408 | 16.4977 | 56152.0802 | 1.127 | 56.1521 | 99.6282 | 5.9215 | 0.099628 | 76.1863 | 4.721 | 0.076186 | <0.000 | N/A | <0.000 | 27.9061 | 3.5215 | 0.027906 | 10.5061 | 87.9691 | 0.010506 | <0.000 | N/A | <0.000 | 20.7002 | 9.6961 | 0.0207 | 2679.107 | 1.3584 | 2.6791 | 198.8218 | 1.4321 | 0.19882 | 162.2233 | 8.3054 | 237112.5433 | 0.49753 | 130626.7867 | 0.55873 | 1936880.48 | 0.18611 | 1332724.7467 | 0.41669 | 2720707.7433 | 0.22625 | 2212005.7467 | 0.52197
| FALSE | 031_CCV.d | 2020-07-16 11:35 | CCV |  | 20 ppb / 1000 ppb | 1 |  | 1106 | 989.9245 | 0.62286 | 989.9245 |  |  |  | 18.9629 | 0.40242 | 18.9629 | 1002.8426 | 0.9319 | 1002.8426 | 965.0953 | 0.55378 | 965.0953 | 20.1745 | 0.19406 | 20.1745 | 20.2169 | 0.089028 | 20.2169 | 999.275 | 1.1501 | 999.275 | 20.653 | 1.191 | 20.653 | 21.0952 | 1.1806 | 21.0952 | 1031.3332 | 0.22817 | 1031.3332 | 20.3689 | 1.1249 | 20.3689 | 20.2773 | 1.1081 | 20.2773 | 20.5713 | 1.8538 | 20.5713 | 18.795 | 0.8079 | 18.795 | 19.745 | 1.5302 | 19.745 | 20.2899 | 0.23854 | 20.2899 | 168.8933 | 8.2153 | 235008.22 | 0.080092 | 129621.3267 | 1.0211 | 1918384.8533 | 1.1069 | 1325901.15 | 0.3706 | 2690127.1867 | 1.1034 | 2178230.1233 | 0.44056
| FALSE | 032_CCV.d | 2020-07-16 11:39 | CCV |  | 1000 ppb (Mg) | 1 |  | 2503 | [9 empty] | 1.0331 | 1316.676 | 1.0331 | [6 empty] | 0.23718 | 13.8266 | 0.23718 |  |  |  | 0.053822 | 10.44 | 0.053822 |  |  |  | 1.7504 | 12.683 | 1.7504 |  |  |  | 0.045913 | 250.9376 | 0.045913 | <0.000 | N/A | <0.000 | 0.058248 | 13.0793 | 0.058248 |  |  |  | 1061.2241 | 14.3476 | 1061.2241 | 146.6667 | 14.1932 | 219335.28 | 12.3059 | 126115.58 | 7.7676 | 1802562.29 | 12.357 |  |  |  |  | 2047948.1833 | 12.3688
| FALSE | 033New2.d | 2020-07-16 11:42 | BLK |  | Cal. Blank QC | 1 |  | 1101 | <0.000 | N/A | <0.000 |  |  |  | <0.000 | N/A | <0.000 | <0.000 | N/A | <0.000 | <0.000 | N/A | <0.000 | 0.0052599 | 34.0566 | 0.0052599 | <0.000 | N/A | <0.000 | <0.000 | N/A | <0.000 | <0.000 | N/A | <0.000 | 0.0029345 | 50.3834 | 0.0029345 | <0.000 | N/A | <0.000 | 0.00076505 | 91.8907 | 0.00076505 | 0.0033704 | 555.2203 | 0.0033704 | 0.1285 | 106.5734 | 0.1285 | 0.016923 | 8.5342 | 0.016923 | 0.0028883 | 132.0009 | 0.0028883 | 0.06103 | 4.7124 | 0.06103 | 178.89 | 19.9243 | 234901.2933 | 1.0577 | 131205.0967 | 0.39745 | 1908457.7033 | 2.0813 | 1327941.33 | 0.86955 | 2705630.5167 | 0.8302 | 2193142.1367 | 1.6768
| FALSE | 034SMPL.d | 2020-07-16 12:06 | Sample |  | 4 | 1000 |  | 2201 | 170727.4333 | 3.1569 | 170.7274 |  |  |  | <0.000 | N/A | <0.000 | 114301.7115 | 1.121 | 114.3017 | 23186.8608 | 4.9285 | 23.1869 | 12.7456 | 8.1516 | 0.012746 | 181.8306 | 1.9121 | 0.18183 | 623.1519 | 15.496 | 0.62315 | <0.000 | N/A | <0.000 | 60.478 | 9.0456 | 0.060478 | <0.000 | N/A | <0.000 | 10.9266 | 12.515 | 0.010927 | 38.5281 | 92.9307 | 0.038528 | <0.000 | N/A | <0.000 | 68.9967 | 11.6108 | 0.068997 | 45.7704 | 10.1122 | 0.04577 | 14.5898 | 2.7701 | 0.01459 | 118.89 | 22.8343 | 235430.4667 | 0.86775 | 132102.74 | 0.26918 | 1931510.41 | 0.64177 | 1330570.0267 | 0.31191 | 2721038.92 | 0.52403 | 2190075.7467 | 0.89468
| FALSE | 035SMPL.d | 2020-07-16 12:10 | Sample |  | 5 | 1000 |  | 2202 | 109098.8505 | 0.2128 | 109.0989 |  |  |  | 9367.9668 | 1.7343 | 9.368 | 396383.0486 | 0.75719 | 396.383 | 233122.1019 | 1.4037 | 233.1221 | 21.9912 | 1.4457 | 0.021991 | 3981.02 | 0.62708 | 3.981 | 15071.5572 | 1.6513 | 15.0716 | 29.1889 | 21.3105 | 0.029189 | 31.5555 | 10.6425 | 0.031556 | <0.000 | N/A | <0.000 | 10.6812 | 8.4025 | 0.010681 | 13.5556 | 199.3666 | 0.013556 | <0.000 | N/A | <0.000 | 19.2091 | 13.8463 | 0.019209 | 480.7581 | 3.3594 | 0.48076 | 4.8623 | 20.7382 | 0.0048623 | 128.89 | 23.3234 | 233537.2133 | 0.41962 | 131236.2433 | 1.518 | 1907629.23 | 0.63182 | 1329791.5133 | 0.34472 | 2694154.4767 | 0.79368 | 2190096.8567 | 0.68674
| FALSE | 036SMPL.d | 2020-07-16 12:13 | Sample |  | 6 | 1000 |  | 2203 | 63252.9512 | 1.645 | 63.253 |  |  |  | 10755.5012 | 0.72518 | 10.7555 | 205764.1428 | 1.5789 | 205.7641 | 117002.4829 | 2.1672 | 117.0025 | 13.7767 | 17.7829 | 0.013777 | 618.3328 | 1.1636 | 0.61833 | 1647.4192 | 1.8589 | 1.6474 | 30.9579 | 16.6075 | 0.030958 | 1733.3791 | 1.2496 | 1.7334 | 354.6772 | 8.7119 | 0.35468 | 6.3017 | 36.3679 | 0.0063017 | 51.0349 | 123.9463 | 0.051035 | <0.000 | N/A | <0.000 | 12.917 | 10.0902 | 0.012917 | 132.5668 | 0.14369 | 0.13257 | <0.000 | N/A | <0.000 | 152.2233 | 7.0396 | 233544.2433 | 0.61197 | 130686.77 | 0.67011 | 1915387.08 | 0.45801 | 1321448.5233 | 0.18064 | 2698344.3367 | 1.1696 | 2186940.5367 | 1.0632
| FALSE | 037SMPL.d | 2020-07-16 12:17 | Sample |  | Method Blank | 1 |  | 2512 | 1.0353 | 21.3299 | 1.0353 |  |  |  | <0.000 | N/A | <0.000 | <0.000 | N/A | <0.000 | <0.000 | N/A | <0.000 | 0.26513 | 1.4075 | 0.26513 | 0.0044591 | 25.4311 | 0.0044591 | 2.9382 | 1.9941 | 2.9382 | 0.062842 | 13.3545 | 0.062842 | 0.071892 | 7.8246 | 0.071892 | <0.000 | N/A | <0.000 | 0.0025691 | 20.5695 | 0.0025691 | 0.024135 | 69.3956 | 0.024135 | <0.000 | N/A | <0.000 | 0.049108 | 10.225 | 0.049108 | 0.020295 | 10.1262 | 0.020295 | 0.0055979 | 10.4733 | 0.0055979 | 163.3367 | 10.2029 | 239148.74 | 0.44188 | 129596.52 | 0.45128 | 1970274.8533 | 0.82609 | 1320769.9567 | 0.21731 | 2722618.5 | 0.8351 | 2202658.6633 | 0.52258
| FALSE | 038_CCV.d | 2020-07-16 12:21 | CCV |  | 20 ppb / 1000 ppb | 1 |  | 1106 | 977.5148 | 0.81347 | 977.5148 |  |  |  | 18.5139 | 1.0156 | 18.5139 | 981.114 | 1.122 | 981.114 | 954.4723 | 0.41638 | 954.4723 | 20.0675 | 0.57579 | 20.0675 | 20.1311 | 0.56391 | 20.1311 | 1001.74 | 1.4178 | 1001.74 | 20.446 | 1.3669 | 20.446 | 20.9373 | 0.79853 | 20.9373 | 1018.2652 | 0.80631 | 1018.2652 | 20.333 | 0.78686 | 20.333 | 20.4001 | 1.1168 | 20.4001 | 20.3462 | 1.9764 | 20.3462 | 18.9532 | 1.189 | 18.9532 | 19.782 | 1.1537 | 19.782 | 19.8986 | 0.36916 | 19.8986 | 171.1167 | 28.9163 | 234467.32 | 0.51457 | 129796.0367 | 0.99917 | 1913484.7867 | 1.4238 | 1326351.1533 | 0.35764 | 2692035.8667 | 1.2876 | 2206170.26 | 0.1017
| FALSE | 039New2.d | 2020-07-16 12:24 | BLK |  | Cal. Blank QC | 1 |  | 1101 | <0.000 | N/A | <0.000 |  |  |  | <0.000 | N/A | <0.000 | <0.000 | N/A | <0.000 | <0.000 | N/A | <0.000 | 0.0086719 | 16.0863 | 0.0086719 | <0.000 | N/A | <0.000 | 0.0023932 | 726.1927 | 0.0023932 | <0.000 | N/A | <0.000 | 0.004298 | 81.8245 | 0.004298 | 0.0012025 | 598.2887 | 0.0012025 | 0.0020895 | 24.6924 | 0.0020895 | 0.01566 | 313.3811 | 0.01566 | <0.000 | N/A | <0.000 | 0.043797 | 10.5383 | 0.043797 | <0.000 | N/A | <0.000 | 0.0029061 | 19.9738 | 0.0029061 | 150 | 17.6383 | 232608.6433 | 0.71343 | 129861.9533 | 0.30082 | 1892500.83 | 1.2179 | 1307907.63 | 0.29129 | 2706518.99 | 0.81764 | 2195855.19 | 1.2344
| FALSE | 040SMPL.d | 2020-07-16 13:49 | Sample |  | 1 | 10 |  | 2201 | [6 empty] | 291.7848 | 0.21531 | 29.1785 | [6 empty] | 23.5245 | 0.47631 | 2.3525 | 1137.3727 | 0.69761 | 113.7373 | 5388.8806 | 1.6335 | 538.8881 | 154.6963 | 1.0792 | 15.4696 | 135.6351 | 0.68545 | 13.5635 | 380.6892 | 0.63672 | 38.0689 | 6.8274 | 2.8588 | 0.68274 | 54.6462 | 2.3671 | 5.4646 | 59.9637 | 1.8773 | 5.9964 | 6.6644 | 1.9069 | 0.66644 | 691.8136 | 1.2662 | 69.1814 | 3.7018 | 2.2373 | 0.37018 | 454.4567 | 9.2871 | 227356.1467 | 0.60272 | 121407.0333 | 0.6113 | 1782536.5967 | 1.5935 | 1190320.22 | 0.43315 | 2596326.4233 | 0.88864 | 2113355.68 | 0.88626
| FALSE | 041SMPL.d | 2020-07-16 13:53 | Sample |  | 2 | 10 |  | 2202 | [6 empty] | 16537.6709 | 1.0515 | 1653.7671 | [6 empty] | 42.3346 | 0.08506 | 4.2335 | 14337.6894 | 0.99878 | 1433.7689 | 10880.4448 | 0.42842 | 1088.0445 | 127.2332 | 0.43708 | 12.7233 | 37.5229 | 1.128 | 3.7523 | 351.7809 | 0.37616 | 35.1781 | 27.3228 | 0.62997 | 2.7323 | 46.0861 | 4.0845 | 4.6086 | 48.5196 | 6.1716 | 4.852 | 16.5646 | 1.8999 | 1.6565 | 2083.6144 | 0.63142 | 208.3614 | 156.9589 | 1.4088 | 15.6959 | 1326.7367 | 1.6478 | 236186.5033 | 0.89368 | 124915.7667 | 0.27057 | 1749444.69 | 0.78247 | 1205601.87 | 0.55576 | 2669575.17 | 0.64409 | 2181448.4567 | 1.0926
| FALSE | 042SMPL.d | 2020-07-16 13:56 | Sample |  | 3 | 10 |  | 2203 | [6 empty] | 67863.6913 | 0.74654 | 6786.3691 | [6 empty] | 76.1842 | 0.15276 | 7.6184 | 17037.0439 | 1.0252 | 1703.7044 | 57595.7143 | 0.59294 | 5759.5714 | 130.9257 | 0.43022 | 13.0926 | 68.0479 | 0.22554 | 6.8048 | 424.2731 | 0.81235 | 42.4273 | 30.7984 | 0.53785 | 3.0798 | 42.9522 | 1.5461 | 4.2952 | 48.5557 | 4.1477 | 4.8556 | 20.5188 | 1.6248 | 2.0519 | 2843.8454 | 1.8209 | 284.3845 | 198.743 | 0.89859 | 19.8743 | 2965.87 | 3.678 | 241882.87 | 0.25298 | 129637.9767 | 0.46504 | 1731735.3133 | 0.13103 | 1207715.01 | 0.51781 | 2717769.06 | 0.76008 | 2204351.51 | 0.57004
| FALSE | 043SMPL.d | 2020-07-16 14:00 | Sample |  | 4 | 10 |  | 2204 | [6 empty] | 51.4763 | 4.1263 | 5.1476 | [6 empty] | 2.6425 | 2.6449 | 0.26425 | 127.2912 | 0.44023 | 12.7291 | 337.4559 | 0.39295 | 33.7456 | 10.4776 | 2.8916 | 1.0478 | 28.4463 | 0.36328 | 2.8446 | 236.4868 | 0.56232 | 23.6487 | 8.7502 | 2.5911 | 0.87502 | 3.1459 | 11.2476 | 0.31459 | 2.1888 | 66.0406 | 0.21888 | 5.1584 | 1.203 | 0.51584 | 24.5304 | 0.40298 | 2.453 | 0.64473 | 1.4584 | 0.064473 | 171.1167 | 17.6757 | 240951.3867 | 0.40646 | 130898.6633 | 0.35188 | 1859534.23 | 0.49404 | 1276382.8933 | 0.28929 | 2668920.7267 | 0.25952 | 2189867.27 | 0.35002
| FALSE | 044SMPL.d | 2020-07-16 14:03 | Sample |  | 5 | 10 |  | 2205 | [6 empty] | 11949.4481 | 0.40999 | 1194.9448 | [6 empty] | 15.0127 | 1.1386 | 1.5013 | 4022.8381 | 0.83796 | 402.2838 | 16211.5655 | 0.42732 | 1621.1566 | 46.9193 | 1.2032 | 4.6919 | 26.1163 | 0.59397 | 2.6116 | 52.5772 | 0.67601 | 5.2577 | 12.26 | 0.86269 | 1.226 | 2.6935 | 26.8355 | 0.26935 | 2.8431 | 60.0345 | 0.28431 | 5.059 | 1.3614 | 0.5059 | 469.7452 | 1.3145 | 46.9745 | 8.8245 | 1.4735 | 0.88245 | 427.79 | 7.1988 | 242031.2467 | 0.66489 | 129743.3567 | 0.32221 | 1836779.0933 | 0.76222 | 1263039.7467 | 0.6227 | 2688229.8233 | 0.96953 | 2193651.0233 | 1.5295
| FALSE | 045SMPL.d | 2020-07-16 14:07 | Sample |  | 6 | 10 |  | 2206 | [6 empty] | 199.8858 | 1.2908 | 19.9886 | [6 empty] | 3.3539 | 0.64913 | 0.33539 | 638.7176 | 0.67309 | 63.8718 | 1921.3419 | 1.6434 | 192.1342 | 19.4564 | 0.19037 | 1.9456 | 34.2944 | 1.3931 | 3.4294 | 59.0665 | 0.64481 | 5.9066 | 6.4327 | 2.0194 | 0.64327 | 5.6645 | 7.9212 | 0.56645 | 4.3718 | 4.4756 | 0.43718 | 5.7494 | 0.76529 | 0.57494 | 112.1592 | 1.4984 | 11.2159 | 2.0247 | 1.4225 | 0.20247 | 241.1167 | 8.9961 | 240675.12 | 0.68501 | 130719.4033 | 0.87509 | 1875042.29 | 0.28766 | 1286184.9467 | 0.37989 | 2704715.8633 | 1.3719 | 2205876.7867 | 0.10895
| FALSE | 046_CCV.d | 2020-07-16 14:11 | CCV |  | 20 ppb / 1000 ppb | 1 |  | 1106 | 1027.6984 | 0.17361 | 1027.6984 |  |  |  | 19.0458 | 0.67984 | 19.0458 | 1055.1508 | 0.73821 | 1055.1508 | 966.548 | 0.31722 | 966.548 | 19.9 | 0.27811 | 19.9 | 20.2197 | 0.36994 | 20.2197 | 1020.0973 | 0.6254 | 1020.0973 | 20.2385 | 1.0722 | 20.2385 | 20.7117 | 0.64865 | 20.7117 | 1036.5297 | 0.22726 | 1036.5297 | 20.1952 | 0.75992 | 20.1952 | 20.6545 | 1.3072 | 20.6545 | 20.3036 | 0.60907 | 20.3036 | 18.3487 | 0.23948 | 18.3487 | 19.4939 | 1.5002 | 19.4939 | 19.5678 | 0.65087 | 19.5678 | 155.56 | 1.2359 | 238524.4 | 0.76941 | 131219.3233 | 0.35868 | 1923235.8967 | 0.25632 | 1312339.2533 | 0.2273 | 2709649.7533 | 1.4681 | 2206192.2067 | 0.37384
| FALSE | 047New2.d | 2020-07-16 14:14 | BLK |  | Cal. Blank QC | 1 |  | 1101 | 2.7177 | 0.29607 | 2.7177 |  |  |  | <0.000 | N/A | <0.000 | 17.6723 | 3.8529 | 17.6723 | <0.000 | N/A | <0.000 | 0.0084798 | 23.4236 | 0.0084798 | 0.0022349 | 126.3426 | 0.0022349 | 0.18062 | 4.5946 | 0.18062 | <0.000 | N/A | <0.000 | 0.0068368 | 31.1815 | 0.0068368 | <0.000 | N/A | <0.000 | 0.0018861 | 15.7545 | 0.0018861 | 0.0029336 | 2829.3947 | 0.0029336 | <0.000 | N/A | <0.000 | 0.067199 | 4.9459 | 0.067199 | <0.000 | N/A | <0.000 | <0.000 | N/A | <0.000 | 168.8933 | 15.9548 | 235362.2733 | 0.35482 | 130638.73 | 0.16648 | 1954019.6433 | 0.59611 | 1302220.3733 | 0.78165 | 2691725.6567 | 0.58748 | 2205779.22 | 0.55247
| FALSE | 048SMPL.d | 2020-07-16 14:18 | Sample |  | 1 | 2 |  | 2207 | [15 empty] | 22.6325 | 0.2607 | 11.3162 | [6 empty] | 151.1832 | 0.14036 | 75.5916 | 137.0054 | 0.10369 | 68.5027 | 377.073 | 0.77705 | 188.5365 | 9.1971 | 1.3773 | 4.5985 | 87.2579 | 0.40846 | 43.6289 | 91.4651 | 0.96913 | 45.7325 | 9.6856 | 2.2108 | 4.8428 |  |  |  | 3.2953 | 0.73913 | 1.6476 | 1542.32 | 2.9427 | 237265.89 | 0.41755 | 120148.9767 | 0.1069 | 1618057.99 | 0.34036 | 1150777.1067 | 0.45049 | 2502873.51 | 0.40852 | 2053116.7233 | 0.78617
| FALSE | 049SMPL.d | 2020-07-16 14:21 | Sample |  | 2 | 2 |  | 2208 | [15 empty] | 39.0993 | 0.21583 | 19.5497 | [6 empty] | 111.9862 | 0.62931 | 55.9931 | 32.0399 | 0.20397 | 16.02 | 319.5456 | 0.23561 | 159.7728 | 33.4559 | 0.63228 | 16.7279 | 65.3946 | 0.78509 | 32.6973 | 68.057 | 1.1294 | 34.0285 | 20.3881 | 2.1418 | 10.194 |  |  |  | 148.1719 | 0.81455 | 74.086 | 4890.8167 | 3.0662 | 243256.2133 | 0.32198 | 128038.34 | 0.57227 | 1465163.0667 | 2.6543 | 1113596.9933 | 0.25131 | 2572080.94 | 0.60162 | 2117879.15 | 1.1712
| FALSE | 050SMPL.d | 2020-07-16 14:25 | Sample |  | 3 | 2 |  | 2209 | [15 empty] | 70.1417 | 0.54505 | 35.0708 | [6 empty] | 105.7339 | 0.59425 | 52.8669 | 54.3374 | 0.29193 | 27.1687 | 350.969 | 0.11438 | 175.4845 | 34.6313 | 0.66141 | 17.3157 | 56.7246 | 0.9704 | 28.3623 | 60.046 | 1.8771 | 30.023 | 24.7901 | 0.97781 | 12.3951 |  |  |  | 174.108 | 0.59602 | 87.054 | 12100.3567 | 1.0494 | 262385.2333 | 0.64131 | 141308.51 | 0.54995 | 1414535.5 | 0.66907 | 1098440.9833 | 0.22607 | 2772484.1967 | 0.55486 | 2268528.8667 | 1.1625
| FALSE | 051SMPL.d | 2020-07-16 14:28 | Sample |  | 4 | 2 |  | 2210 | [15 empty] | 2.5622 | 1.4642 | 1.2811 | [6 empty] | 10.7848 | 1.0528 | 5.3924 | 27.9186 | 0.67651 | 13.9593 | 241.3482 | 0.32475 | 120.6741 | 9.3977 | 1.012 | 4.6988 | 3.5897 | 6.3893 | 1.7949 | 3.4459 | 12.367 | 1.7229 | 5.4757 | 0.50518 | 2.7378 |  |  |  | 0.7479 | 1.1382 | 0.37395 | 304.4533 | 7.2902 | 250158.9733 | 0.91649 | 130025.4233 | 0.8493 | 1813240.5533 | 1.2142 | 1244548.03 | 0.17057 | 2572554.0633 | 0.71036 | 2132314.9133 | 0.78547
| FALSE | 052SMPL.d | 2020-07-16 14:32 | Sample |  | 5 | 2 |  | 2211 | [15 empty] | 14.1474 | 0.71431 | 7.0737 | [6 empty] | 44.5735 | 0.74194 | 22.2867 | 24.3028 | 0.91766 | 12.1514 | 53.7232 | 0.74506 | 26.8616 | 12.4773 | 0.64177 | 6.2387 | 2.9334 | 6.0872 | 1.4667 | 3.5545 | 7.2787 | 1.7772 | 5.4861 | 0.90764 | 2.7431 |  |  |  | 8.4584 | 0.37724 | 4.2292 | 1534.5367 | 4.314 | 257750.8767 | 0.57228 | 132099.29 | 0.57255 | 1701653.06 | 0.87861 | 1192443.8933 | 0.38766 | 2613132.3967 | 0.17855 | 2124911.65 | 0.11641
| FALSE | 053SMPL.d | 2020-07-16 14:35 | Sample |  | 6 | 2 |  | 2212 | [15 empty] | 3.2675 | 0.27334 | 1.6337 | [6 empty] | 19.9064 | 0.86156 | 9.9532 | 32.8897 | 0.35529 | 16.4449 | 63.0091 | 0.8978 | 31.5046 | 7.0749 | 1.0596 | 3.5375 | 6.4301 | 0.99729 | 3.2151 | 6.7899 | 6.2264 | 3.395 | 6.0938 | 1.8066 | 3.0469 |  |  |  | 2.0372 | 0.55155 | 1.0186 | 481.1233 | 11.2647 | 249237.16 | 0.42593 | 128651.4433 | 0.70004 | 1794657.0833 | 2.4699 | 1233115.6633 | 0.52127 | 2606314.7567 | 0.52735 | 2140966.2367 | 0.15305
| FALSE | 054_CCV.d | 2020-07-16 14:39 | CCV |  | 20 ppb / 1000 ppb | 1 |  | 1106 | 1065.0651 | 1.0582 | 1065.0651 |  |  |  | 19.4771 | 0.93356 | 19.4771 | 1275.031 | 0.44526 | 1275.031 | 972.0357 | 0.14332 | 972.0357 | 19.5618 | 0.91276 | 19.5618 | 20.0925 | 0.40266 | 20.0925 | 1020.7078 | 0.83586 | 1020.7078 | 20.0131 | 0.12217 | 20.0131 | 20.301 | 0.17474 | 20.301 | 1022.9684 | 0.48002 | 1022.9684 | 20.1481 | 0.52985 | 20.1481 | 20.5308 | 1.3547 | 20.5308 | 20.5507 | 2.1944 | 20.5507 | 18.4332 | 0.64259 | 18.4332 | 19.9862 | 0.26002 | 19.9862 | 19.0794 | 0.88268 | 19.0794 | 188.89 | 6.1965 | 250443.8867 | 0.75271 | 135247.3933 | 0.36183 | 1937572.0067 | 0.41282 | 1337201.3167 | 0.74513 | 2698102.3233 | 0.30543 | 2211531.6467 | 0.75062
| FALSE | 055New2.d | 2020-07-16 14:42 | BLK |  | Cal. Blank QC | 1 |  | 1101 | 32.5757 | 3.6193 | 32.5757 |  |  |  | <0.000 | N/A | <0.000 | 193.4278 | 1.891 | 193.4278 | <0.000 | N/A | <0.000 | <0.000 | N/A | <0.000 | 0.02379 | 5.2506 | 0.02379 | 0.35622 | 8.139 | 0.35622 | <0.000 | N/A | <0.000 | 0.0050649 | 28.6331 | 0.0050649 | 0.022479 | 103.7569 | 0.022479 | 0.0042842 | 18.6177 | 0.0042842 | 0.10833 | 38.663 | 0.10833 | <0.000 | N/A | <0.000 | 0.085897 | 9.165 | 0.085897 | 0.0013548 | 181.8725 | 0.0013548 | <0.000 | N/A | <0.000 | 170.0033 | 15.3122 | 245825.2333 | 1.0272 | 133625.21 | 0.31908 | 1940728.4633 | 2.0206 | 1319970.5433 | 0.24043 | 2685375.24 | 0.14199 | 2197664.9833 | 1.0139

## ValueList_Helper (hidden)
Bkgnd
BlkVrfy
CalBlk
CalStd
CICSpike
DilStd
DriftChk
FQBlk
IsoStd
QCS (QC1)
LFB (QC2)
ICV (QC3)
CCV (QC4)
QC5
Sample
Spike
Spike Ref
SQBlk
SQISTD
SQStd
BLK
CCV
DUP
ICV
LFB
LFBD
LRB
LRBD
LRBS
QCS
Spike QC
SpikeQC2